Correlation between staggered BPM and RF-BPM readings:

Calibration of their (linear) dependence and reliability of it

Thomas Wehrli, Abt.4n (ETH)

Abstract

For the 07 beamline at Swiss Light Source, we calculated calibration factors for both XBPMs once by applying a symmetric bump at the BPMs and once with an anti-symmetric. The theoretical readings at the XBPMs were determined under consideration of optical influences that change all calibration factors about 10%.

With that, we got a calibration factor of about 0.95 (symmetric) and 1.06 (anti-symmetric) for the first XBPM and 0.77 (symmetric) and 0.84 (anti-symmetric) for the second XBPM.

As the symmetric and the anti-symmetric values are not consistent within their statistical errors of less than 0.1%, some relevant unconsidered systematic effects are described here. The biggest of them is more than 30 times bigger than the rest of our uncontrolled inaccuracies and noise, which limit the beam stability. An elimination of this effect therefore has a potential not only to make the calibration factors more consistent, but even to reduce the root-mean-square deviation of the beam position.

1 Introduction	3

2 Theoretical background and experimental setup	3

2.1 How to fix the beam	3

2.2 Calibration factor	4

Influence of the Optics	5

Further influences and their consequences	6

2.3 Motivation for the fit curves used	6

3 Observations and Results	6

3.1 Alignment of the front-end diaphragm	6

Situation after our alignment – and before	7

3.2 Calibration factors from different Taylor fits	9

3.3 Further fits	11

4 Discussion	12

4.1 Elimination of systematic effects	13

Further conclusions	15

5 Outlook	15

Literature	16

Glossary	17

Acknowledgement – Dankansagung	18

# 1 Introduction

In many cases, the most successful approach to get more information about chemical structures is photon diffraction. To do this as good as possible, one needs a monochromatic (X-ray) beam that is as intense as possible and constant in many ways. – For us, this means in particular that the beam should be fixed in space on the target.

Today, the best (financially possible) way to get the wanted radiation is to build circular particle accelerator for electrons. The charged particles, with a kinetic energy of 2.4GeV in our Synchrotron, deliver the highly intense beam. (More about this e.g. in [1])

To get the wanted constancy of the beam in space, which should be better than ±1μm, many parameters have a certain influence. For example, resonances of the solid and heavy girders, on which the different magnets are mounted, play a role. They can make the experimental set-up vibrate in bad cases up to microns resulting in a few microns of beam vibration. At the SLS, girders move at least one order of magnitude less, but still are one little parameter among others that contribute to the beam movement.

As we have many more and much bigger influences such as the mains and the insertion devices, the maximum we can reach with static corrections (e.g. better girders) is beam stability on the micron level. Therefore, we constantly have to correct the drifts and vibrations of the beam to get a dynamic equilibrium that fulfills our wanted precision better than ±1μm.

In this report, we consider about the dynamic corrections, which are performed by a fast orbit feedback (FOFB). In particular, we analyze the XBPMs, which measure a beam position changing. With that information, the FOFB can determine the needed BPM positions to bring the beam back to the wanted reference.

# 2 Theoretical background and experimental setup

## 2.1 How to fix the beam

Due to the fact that a manipulation of charged particles is easier than for photons, the photon beam position is best changed via the electron orbit. Two BPMs make it is possible to deflect the orbit with a symmetric bump or to realize an angle with an anti-symmetric bump. During experiments, the BPMs are steered by the fast orbit feedback, which has its information about the deflection of the photon beam position compared to a reference from two staggered BPMs (see fig.1).

Fig. 1: Schematic overview of the experimental setup for a symmetric bump. The picture is not length preserving, the most relevant real distances are 1.4m from the source point to each BPM and 4.1m to the 1st XBPM and 6.1m to the 2nd.

## 2.2 Calibration factor

The XBPM blades deliver a photoelectric current of only about 100μA, depending on the photon flux at them. This signal is – although it is extremely small – almost noiseless. According to the producer of the involved Low Current asymmetry detector, noise is less than 0.1% of the signal in this range. This makes it possible to register extremely small changes of the photon beam position. Nevertheless, due to many small unseizable effects, it is not possible to calculate the absolute position of the photon beam quantitatively exactly from the blades, so that we have to calibrate the readings. Therefore, we drive quantitatively well-defined symmetric and anti-symmetric bumps at the BPMs. These bumps cause theoretical readings at the XBPMs, which are compared to the measurement.

To get one measured reading from the four blades per XBPM, we build asymmetries

asym1 = (b1–b3) / (b1+b3), asym2 = (b2–b4) / (b2+b4), (b1,…, b4 the blades, numbered according to fig.2), which are transformed to expected distances P1 and P2 in "millimeters", according to

(1).

As these positions must be equal, we get

(2)

Here, c is what we call experimental XBPM reading or, actually more correct, SPM reading. The meaning of the offset becomes clear from fig.2.

Fig.2: Schematic view of a staggered BPM. The upper left blade is b1, the upper right b2, the lower left b3 and the lower right b4. The upper (lower) dash-dotted line marks the center of the left (right) blade pair. (Courtesy Juraj Krempasky)

The theoretical XBPM reading can be determined from the quantitatively well defined BPM bump B. With that, the calibration K becomes

(3)

The factor I transforms the applied BPM bump to an expected XBPM reading; it depends on the bump type as well as on geometrical and optical factors.

If we draw I B on the x-axis against c, the calibration factor is given by the slope S of our curve.

(4)

(Remark:

The approach of equation (1) uses a linear dependence of the positions P1 and P2 from their according asymmetry. As this is a simplification, the real position of the beam c and the single-asymmetry positions P1 and P2 lose their direct proportionality and we expect that c goes faster than linearly towards ±∞ for big bumps. This is due to the asymptotic behaviour of the asymmetries and leads to an artificial reduction of the calibration factor if the beam leaves the center of the blades.)

### Influence of the Optics

Because of the quadrupoles between the BPMs, the theoretical XBPM reading I B cannot be calculated just from geometric factors and the BPM bump according to fig.1, as all magnetic components have an optical influence on the beam.

The quantitative influence of the optics must be calculated with a model, as we only have measured data at the BPM positions; fig.3 shows such a calculation.

Fig. 3: Calculated deflection for a 1mm symmetric bump. The thin arrows mark the BPMs and the bold ones the position before and after the dipole magnet.

(*As our model is linear in good approximation, the figure is valid for any length unit in the range below mm.)

According to fig.3, we can calculate the factor I that transforms the BPM bump B to an expected XBPM reading IB for both XBPMs and bump types:

(5a)

(5b)

(5c)

(5d)

(rbef is the RF-BPM reading before and raft after the dipole magnet, dBPM the distance between the two BPMs, d1 and d2 the distances from the source point to the 1st and to the 2nd XBPM. rsp = (rbef + raft) / 2 is the theoretical reading at the source point, which is in the middle of the dipole magnet.)

The purely geometric values for I would be 1 for (5a) and (5c), 2.93 for (5b) and 4.36 for (5d). Hence, the optical influence enhances all calibration factors a bit less than 10%.

### Further influences and their consequences

The most disturbing influences come from shadowing effects. These can even cause a not injective behaviour of the blades’ asymmetries in worst case, which makes it impossible to assign a certain reading to one deflection. This makes a linear calibration factor lose all of its meaning. Shadowing can be reduced by moving the responsible parts of the measurement installation, namely the front-end diaphragm or one of the XBPMs.

Further challenges come from the fact, that actually every component of the experimental setup and every parameter might have a certain influence on the XBPM readings. To describe their influences theoretically is a big task.

A little example for such a task is the influence of the voltage impressed on the blades, which is done to avoid that electrons can jump from one blade to another. Some questions here are, how many electrons can jump as a function of our voltage? What does that mean if the blades are not equal because of any other reason? Does a high voltage change any other parameter?

As these are not the only questions, it becomes plausible, that it is difficult to find an exact description for our whole XBPM readings.

## 2.3 Motivation for the fit curves used

As described above, different aspects make it difficult or even impossible to calculate a theoretical fit curve that describes our XBPM readings in an – apart from noise – exact or at least helpful way.

The most obvious approach in this case is to use a polynomial fit. As long as the data points can be fitted with a smooth curve, we get a Taylor approximation, which is good enough for our purposes; in fact, the main thing we need in the end is just the slope of the curve around zero.

The order of the polynomial fit we used was usually three. We didn’t use a higher order because a simultaneous estimation of more parameters that are even not physical is often ambiguous, especially with only 40 or 80 data points. Depending on the personal preferences, statisticians say that 5-6 or at least 10 data points per parameter are necessary, but some even add per physical parameter.

In this sense, we tried to add as much physics as we can. – As the beam profile is symmetric in first order, the SPM readings should have an odd symmetry around zero if the blades are aligned exactly around that point. Of course, this is not possible perfectly and thus, the symmetry is not exactly around zero. Therefore we tried a(x- x0)3 + b(x- x0) + c. Here we also have four parameters, but x0 has a physical meaning (misalignment of the blades). This allows us to control whether the fit makes sense in a physical way or not. A further step on this way is to estimate x0 with additional criteria so that we can fix it before fitting. With that, we reduce the number of parameters from four to three.

# 3 Observations and Results

## 3.1 Alignment of the front-end diaphragm

As we had a bad illumination of the 2nd XBPM because of shadowing effects, we had to shift the diaphragm towards the inner of the storage ring. Fig.4 shows the changing we did.

Fig.4: Dipole chamber “CBX07”. With the marked screws, the diaphragm was shifted about 1.5mm towards the inner of the storage ring. The positions of the screws are: Screw 1: 16.0 (before shift) / 14.5 (after); Screw 2: 16.7 / 15.3; Screw 3: 16.4 / 14.6 [mm]. (Courtesy Lothar Schulz)

### Situation after our alignment – and before

The illumination of our 2nd XBPM became much better with our alignment, so that we got rid of the very harmful not bijective behaviour of the reading (see fig.5).

Fig.5: Readings of the 2nd XBPM before the alignment (left) and afterwards (right).

The 1st XBPM reading showed a bijective behaviour before and after the alignment and the illumination didn’t change obviously. Nevertheless, a little changing could be seen by applying horizontal bumps. – As our radiation describes a quite broad fan beam, such a bump shouldn’t change the readings at the SPMs, because they only detect vertical profiles. If the sum signal nevertheless is a function of the horizontal bump, this is an indication that the fan beam is not aligned around the center of the blades horizontally very well. A further indication for such a misalignment is that the sum signal becomes smaller the bigger the misalignment is. Fig.6 shows that the sum signal before the shift indeed changed a bit.

Fig.6: Sum signal as a function of the 1st BPM reading before the alignment. The arrows show the chronological way of the data acquisition.

After the shift, we got behaviour like in fig.7.

Fig.7: Sum signal as a function of the 1st BPM reading after the alignment. The arrows show the chronological way of the data acquisition during the jumps. (To keep a better overview, the rest of the obvious way is not shown with arrows.)

The first thing to mention about fig.6+7 is that we get some jumps. – These are an unknown influence on the experiment that should be avoided (if possible) or considered and corrected. However, the important properties of the readings can be seen: Before the shift, we get a sum signal difference of ~6μA for the anti-symmetric and ~4.5μA for the symmetric bump on about 1000μA. Afterwards we only get about two thirds of those differences – though with only a third of the former bump amplitude! In fact, we get a difference that is twice as big as before. Additionally, the total sum signal is about 1% smaller. This could mean that we pay the much better behavior of the 2nd XBPM with a slightly worse of the 1st. (The possibility that all we see here could come from unconsidered effects is discussed in section 4.1, Further conclusions.)

## 3.2 Calibration factors from different Taylor fits

Taylor approximations were done in first and in third order for both XBPMs and both bump types. From these fits, we derived the slope of the fit curve in or around zero and from this the calibration factor in the respective point, according to equation (4). Table 1 gives a numeric overview of the results for the 1st XBPM, table 2 for the 2nd one. In fig.8+9, the overview is shown graphically. The difference plots in the right column are the graphical equivalent to the error bars, but with the advantage that they show systematic differences, too. Thus, one can see systematic deviations with even characteristic for the second XBPM that become relevant for 0.3mm bumps. In all other cases, the 3rd order Taylor fits don’t show obvious and big systematic deviations or only with mainly odd characteristic, which is compatible with the assumption that our XBPM reading has an odd symmetry.

Fig. 8: Measured 1st XBPM reading against theoretical reading IB for anti-symmetric bumps (upper half) and symmetric bumps (lower half). The plots in the right column are difference plots between fit curves and data points of the according diagram on the left side. (Note the different scale of the linear and cubic difference plot.)

| fit curve, evaluation point symmetric bump (anti-symmetric bump) | Symmetric bump | Anti-symmetric bump |
| --- | --- | --- |
|  | 0.1mm | 0.3mm | 0.1mm | 0.3mm |
| linear | 0.941 ± 0.002 | 0.914 ± 0.005 | 1.032 ± 0.004 | 0.839 ± 0.012 |
| cubic, 10μm (3.4μm) | 0.9540±0.0014 | 0.9538±0.0008 | 1.0662±0.0006 | 1.0754±0.0010 |
| cubic, 0μm | 0.9487±0.0012 | 0.9491±0.0007 | 1.0616±0.0006 | 1.0711±0.0009 |
| cubic, -10μm (-3.4μm) | 0.9428±0.0014 | 0.9440±0.0008 | 1.0567±0.0006 | 1.0662±0.0010 |

Table 1: Calibration factors for the 1st XBPM and different evaluation procedures. The 95% confidence intervals come from Origin 7.5 estimations of the relevant parameters.

Fig. 9: Measured 2nd XBPM reading against theoretical reading IB for anti-symmetric bumps (upper half) and symmetric bumps (lower half). The plots in the right column are difference plots between fit curves and data points of the according diagram on the left side. (Note the different scale of the linear and cubic difference plot.)

| fit curve, evaluation point symmetric bump (anti-symmetric bump) | Symmetric bump | Anti-symmetric bump |
| --- | --- | --- |
|  | 0.1mm | 0.3mm | 0.1mm | 0.3mm |
| Linear | 0.7618±0.0021 | 0.752 ± 0.005 | 0.804 ± 0.007 | 0.617 ± 0.017 |
| cubic, 10μm (2.3μm) | 0.7746±0.0017 | 0.7723±0.0007 | 0.8347±0.0004 | 0.853 ± 0.007 |
| cubic, 0μm | 0.7680±0.0015 | 0.7665±0.0007 | 0.8315±0.0004 | 0.846 ± 0.007 |
| cubic, -10μm (-2.3μm) | 0.7610±0.0017 | 0.7607±0.0007 | 0.8239±0.0004 | 0.838 ± 0.007 |

Table 2: Calibration factors for the 2nd XBPM and different evaluation procedures. The 95% confidence intervals come from Origin 7.5 estimations of the relevant parameters.

## 3.3 Further fits

The programs we used to fit the data (Origin 7.5. and different! Excel versions) produced results for the approach with y = a(x- x0)3 + b(x- x0) + c (a, b, c, x0: constants) like in fig.10.

Fig.10: 1st XBPM reading and fit curve with y = a (x - x0)3 + b (x - x0) + c (a = 0.478; b = 0.909; c = 0.204;

x0 = 0.131) for an anti-symmetric bump (left); difference between fit and data (right).

This fit is – within round off errors – equal to the according Taylor fit, which is nothing but normal as

a(x - x0)3 + b(x - x0) + c = a’x3 + b’x2 + c’x + d’

with                  a′ = a, b′ =-3 ax0, c′ = 3 ax02 + b, d′ = c - bx0 - x03    and x0 = - b′ / 3a                        (6)

is valid. Nevertheless, it took a big effort with Excel to get this result. Depending on the starting values and the number of iteration steps, local minima effects and other influences cause much different results with x0 values of almost 5mm. With our physical criterion of x0 (values >1mm are nonsense), it was at least directly visible that something was wrong with our fit. However, one has to keep in mind that fitting problems always can occur and only some of them are obvious. This is one reason why we should add as much physics to our fit curve as we can.

For example, we could previously estimate x0 with the sum signal of all blades. The idea is based on the fact that the sum signal is minimal if the beam is exactly in the middle of the XBPM. This is – for equal blades – true as long as the vertical beam profile in the detected ranges falls faster than 1/x, which is given for the ends of all possible beam profiles, as their integral over the whole space must be finite. With that and as we only detect the ends of the beam, we can estimate x0 from the sum signal (see fig.11) and enter this in our approach.

Fig.11: Sum signal for a vertical, anti-symmetric bump and quadratic fit curve with y = (2.85x2 – 0.32x + 1.05)·103.

An advantage of this idea is that we have to determine one parameter less in our fit curve as x0 is fixed now. A certain problem is of course that the estimation of x0 is not exact. In fact, we got estimations for x0 from different fits and estimations between 166μm and 183μm for anti-symmetric bumps (i.e. <x0> = 175±8 μm) and between 164μm and 188μm for symmetric bumps (<x0> = 176±11 μm). At least, one can say that the symmetric and the anti-symmetric estimations are completely consistent within the statistical reliability of the data.

The fit with our fixed x0 = 175μm is depicted in fig.12.

Fig.12: 1st XBPM reading for an anti-symmetric bump and fit curve with y = a (x-0.175)3 + b (x- 0.175) + c (a = 0.413, b = 0.932, c = 0.259,) for an anti-symmetric bump (left); difference between data and fit (right).

A comparison of fig.12 with fig.11 shows that the differences between fit curve and data become one order of magnitude bigger by fixing x0. Unchanged is the trend to bigger differences if IB goes towards one. This is in opposite to the linear fit of the same data (see fig.8) where the biggest differences, which are one order of magnitude bigger again, occur if IB goes towards minus one.

# 4 Discussion

One question raised in this report is how the data should be evaluated and interpreted.

The reason why we didn’t use the approach with the fixed x0 = 175μm to get our calibration factor is that the Taylor approximations showed much less deviations from the data. The price is of course that the x0 values recalculated according to equation (6) can differ a lot from the values we got from the sum signal. E.g., for the 1st XBPM, 0.1mm symmetric bump, we got x0 = 93μm instead of 175μm. We accepted this because the value of 175μm can be systematically wrong, too, e.g. if the blades were not equal.

Another point in this context is the interpretation of the error bars from Origin 7.5. They are based on the wrong assumption that we only have statistical deviations between data and fit curve. So, mainly they tell us that we have systematic deviations, which are orders bigger than the statistical ones.

Even if we got rid of all fit curve based errors, there would still be systematic errors making the confidence intervals like those from Origin doubtable. Looking at the last ones of them, which come from the real move of the beam (the reason why we do all of this!), we see that they cause an error that is a bit bigger than Origin’s values yet.

The beam movement can be seen for example if one plots the difference of the two values for the same BPM position and both XBPMs (see fig.13).

Fig.13: Difference Δ between the first and the second data point for a symmetric BPM bump for the 1st (left) and 2nd (middle) XBPM. The difference between the two Δ-values is depicted right. (Note the scale.)

*The 1st XBPM reading is calibrated with a calibration factor of 1, the 2nd with 0.8.

These small effects of course don’t really harm our calibration factor; they only change it in the low per mil range. In fact, our calibration constant just has to give a rough order of the dependence between BPMs and XBPMs anyway. After all, a not exact factor simply makes the fast orbit feedback correct 18% or 22% instead of the actually wanted 20% of the beam’s zero deviation per feedback loop. In that sense, it is not problematic at all to use a calibration factor of 1 for the 1st XBPM and 0.8 for the second.

If the determination of the calibration factors was the only purpose, we could even afford the luxury of not finding the systematic effect, which makes the symmetric and anti-symmetric calibration factors 10% different. Nevertheless, it would not be wrong of course to check possible reasons such as the influence of the optics or in general the calculation of the theoretical XBPM reading again.

## 4.1 Elimination of systematic effects

In the discussion above, we say that our calibration factors meet all requirements of us. This is not wrong, but one main reason why we did not just make an auto-calibration with our setup was that we wanted to find and exclude all effects with the potential to harm our whole measurement. As we still have systematic effects we didn’t assign to a reason and correct, it shall be outlined here what one further can do against this.

One approach to eliminate systematic effects is to calibrate every single blade. – An advantage of single blade fits is that it becomes easier to calibrate the XBPMs with a physical fit curve. (The only thing one has to do is to determine the beam profile and to assume a certain relation between the photon flux at the blades and their current. With that, it is straightforward to find the suitable fit curve.)

An even bigger advantage is that we can recognize and correct systematic effects, in particular shadowing and signal jumps, for all blades individually. In our case, the first-mentioned could help to explain why we have a certain even behaviour of the readings, in particular of the 2nd XBPM for big bumps. The second-mentioned can be seen in fig.14.

Fig. 14: Current of the four blades as a function of the time for no bump. The vertical lines (left) mark the positions, where the current jumps upwards for all blades. This artefact comes from the convolution of our reading with a sawtooth function. On the right side, the rough trend of the real reading from 100s on is shown.

We see that the changing of the blades’ current in the range depicted in fig.14 is dominated by the convolution of the signal with a sawtooth function. That’s why all four blades show a similar behaviour on the first view. If the currents indeed change because of a movement of the beam, it is given that blade one and two do the opposite of three and four.

The long-term trend of the real – for blade two and four opposite – signal that sits on our sawtooth, is indicated on the right side of fig.14. Here, we see that the blades can register movements much below the order of the sawtooth’ effect.

Remark: The sawtooth periods are not exactly equidistant. In a time range of 1200s where no bump was applied, we had periods between 167s and 179s; the jump duration was 3-6s. With bumps, the periods can differ even more with these values (see fig.15). This makes it impossible to deconvolute our signal in previous as long as we don’t have further information about the sawtooth.

A way to deal with the sawtooth’ artefact is that we can get a beam position from every single blade. With these four results, we can easily see that the reliability of our determined beam position is different as a function of time. If we build the mean <P> with our four positions P1,…, P4, a big changing of the standard deviation of this tells us where we have our jumps. At these positions, our fast orbit feedback probably shouldn’t trust to the XBPM information too much. On the other hand, for all other positions, the deconvoluted readings might be very consistent. Therefore, we could correct more than 20% of the zero deviation.

What exactly a new correction mechanism should look like must be shown in further investigations. There we could also think for example to correct long-term movement in previous, comparable to a feed forward mechanism that is used in undulator setups.

### Further conclusions

A further effect of the sawtooth can be seen in fig.7. The quite chaotic behaviour of the sum signal comes from it, too. If we deconvolute the anti-symmetric sum signal, we get a signal that changes about 2μA on a 0.1mm anti-symmetric bump (see fig.15), which is, related to the bump, the same as the 6μA on 0.3mm before the shift (see fig.6). Additionally we get rid of the unexplained hysteresis before the shift. This could mean that the 1st XBPM behaves rather more convenient after the alignment, too. However, the total signal is smaller after the shift.

Fig.15: Deconvoluted anti-symmetric sum signal after the alignment of the diaphragm (compare fig.7). We assumed a sum signal decay of 51nAs-1 and a jump of 3250nA, which gives a sawtooth period of only 64s, compared with ~170s from fig.14.

Remark: The point at -0.05mm during the jump was corrected to a reasonable value by hand.

The situation for the symmetric sum signal is comparable, so that we waive of an explicit depiction of it.

# 5 Outlook

The higher question in the field we treated here is how good can we fix the beam? As the limitation of the beam stability is given by the inaccuracy of the blades’ current, we have to estimate this reliability: From fig.14, we get a current jump ΔI of about 500nA for blade 1. These jumps are one source of inaccuracy we don’t control so far. The distribution mainly caused by further sources can be seen if we plot one blade’s current against another (see fig.16) .

We can see from fig.16 that the total noise current for blade 1+2is in the order of 30nA. This gives a noise current of about 13nA for blade 1 and 27nA for blade two if we distribute it proportionally., which is about a factor of 37 smaller than the sawtooth jump (blade 1 ~500nA, blade 2 ~1μA). Further investigations have to show, how much of this factor is averaged out by constructing the XBPM reading c. So far, we only can assume that the sawtooth leads to a relevant part of the unwanted beam movement. This is in particular because the handful quite wrong XBPM values from it forbid that we correct more than 20% of the beam’s zero deviation at the moment.

Fig. 16: Dependence of two blade currents and linear fit (right). To get the graph on the left, the differences between the linear fit and the 1196 data points were built. The resulting 1196 values were divided into 1nA classes and plotted against their frequency.

A further step on the way to even better beam stability could begin with an analysis of the shape of fig.16, left side. It is quite probable to find further effects causing this.

Hence, further steps to improve our correction mechanism are still open; the goal to get a beam movement that is a fraction of the actual one might be – at least – possible.

# Literature

[1]	H. Wiedmann, Particle Accelerator Physics I, Basic Principles and Linear Beam Dynamics, Springer Verlag

# Glossary

Bending (magnet): Magnet in the storage ring to make electrons moving on a circle.

Booster (magnet): See Linac

BPM: Beam Position Monitor. The abbreviation is used for the electron beam position monitors (See also RF-BPM; XBPM, staggered BPM)

Fast orbit feedback (FOFB): Correction system that uses the XBPM information to centre the photon beam.

Feed forward: Mechanism used to correct the systematic effect, which would be caused by a changing of an undulator’s gap, in previous.

Front end: Part between the storage ring and the beamline. The XBPMs are located there.

ID: Insertion device. Among other things, device, with which one produces an intense photon beam. There are two types of IDs, undulators and wigglers, which give beams with different characteristics.

Linac: A linear accelerator delivers preaccelerated electrons (100MeV) to a booster that adds 2.3GeV to the electrons, which are then injected into the storage ring with the wanted energy of 2.4GeV.

Optic: Optical components modify the beam. – In (photon) optic, these are e.g. mirrors. If the beam consists of electrons, the optic is given by the electromagnetic conditions. So, for example all magnets are designated as “optical components”, as they influence the beam.

RF-BPM (see also BPM): Instrument to control electron beam position in the storage ring. As the cycling frequency of the electrons is in the radio frequency range, the according BPMs are called RF-BPM (or only BPM).

Source point: origin of the photon beam

SPM: Staggered BPM

Staggered BPM: Staggered BPMs are used for photon beam control of dipole beamlines (see also XBPM). The staggered form is most appropriate if the horizontal beam profile is not considered. As the XBPM is the known abbreviation for a photon BPM we use it here also for our staggered BPMs.

Top-up operation: Operation mode of a Synchrotron light source where electrons are injected into the storage ring in time intervals of only few minutes to keep the current and with that the photon beam intensity more constant. – At the SLS, we inject as soon as the current has decayed about 1mA on a total storage ring current of 350mA. (If one operates with the other used mode (decaying beam mode) one waits for hours between two electron injections.)

XBPM: Beam Position Monitor (BPM) to control the position of the photon beam in an undulator or wiggler beamline. The X stands for the form of this kind of BPM. The abbreviation is also used for all photon BPMs.

# Acknowledgement – Dankansagung

First, I want to thank everyone with whom I have worked that he spoke my language, which is on the one hand the language of a not knowing student and on the other hand German. I hope that the communication was fine in both directions, once through this report and then through my “German”. (I’m Swiss.)

As my written German might not be a linguistic problem to anyone who had to do with me during my time at PSI, I will have the acknowledgment in that language.

Vorgeschlagen wurde mir dieses Praktikum, nach persönlicher Beratung, von Christoph Grab (ETH). Nicht selbstverständlich ist im Weiteren, dass Leonid Rivkin (PSI) tatsächlich bereit war, für einen Studenten eigene Zeit und die seiner Mitarbeiter herzugeben. Für die interessante Praktikummöglichkeit und die dafür investierte Zeit danke ich ihnen freundlich.

Ein besonderer Dank geht natürlich an Michael Böge und Juraj Krempasky für die Betreuung während meiner Zeit am PSI. Ich danke, dass sie trotz anderer Verpflichtungen praktisch immer Zeit für mich fanden und sogar aufforderten, einfach "auf den Wecker" zu gehen. Insbesondere danke ich auch für die lehrreichen Verbesserungsvorschläge und die schonungslos konstruktive Kritik zur ersten Fassung dieser Arbeit.

Bedanken möchte ich mich natürlich auch bei allen andern, die sich mit mir abgaben. So geht ein Dank unter anderem namentlich an Jörg Raabe für die gleichmütige Beantwortung meiner – so vermute ich – kaum immer intelligenten Fragen und für weitere wertvolle Hinweise; an Rolf Wullschleger für die Erläuterungen zur Technik und des gesamten experimentellen Aufbaus; an Thomas Schmidt für die Einführung in "feed forward"; an Lothar Schulz für das Foto der Dipol-Kammer und an Colin Higgs dafür, dass er alle benötigten Programme zum Laufen brachte.

Fürs Teilen des Büros und die angenehme Gesellschaft und nicht zuletzt fürs Organisieren eines Spinds geht ein freundlicher Dank an Christine Kunz.

Als vorletztes möchte ich noch Petrus erwähnen, der durch seine teilweise speziellen Wetterverhältnisse abenteuerliche Fahrradfahrten ans PSI ermöglichte. Ich danke für diese aufregende Zeit, nicht nur auf dem Fahrrad.

Ein grösster Dank geht aber natürlich an meine Eltern dafür, dass sie für die Bahnkosten – und nicht nur das – aufkamen, wenn sich Petrus wieder einmal etwas gar Spezielles einfallen liess.

Danke

Hedingen im Januar 2006 	Thomas Wehrli

# Footnotes

We usually will use the abbreviation "XBPM" for the staggered BPMs too, as this is the known short cut for a photon BPM from undulator setups (see also Glossary). Also in use here is “SPM”.

MMS Frank Optic Products GmbH, Rudowe Chausee 29, 12489 Berlin; www.GMS.Teleport-Berlin.de

To signalize the difference between real and expected distances we will denote the second ones in quotation marks (e.g. “mm”).

See fig.8, anti-symmetric bumps, 2nd plot

Further values for x0 were 143μm (anti-symmetric 0.1mm bump), 131μm (anti-symmetric 0.3mm bump) and 140μm (symmetric 0.3mm bump).

We began with bump deflections of -a, drove until +a in equal steps and went back to -a. This gives two points for every position except +a.

Long-term means something between 100s and 1000s in this case.

Such information could come from the origin of the sawtooth. – It seems plausible that the top-up operation of the SLS is responsible for it.  Every ~3min roughly, we inject about 1mA on a total of 350mA into our storage ring. Further investigations have to show if this approach leads to useful information.

Blade 1+2 are both above the beam and therefore show a linear dependence from each other in the not deconvoluted state yet. As a deconvolution is not simple (e.g. because of the not equidistant jump periods) we refrained from it and accepted the relatively few bad values produced during the jumps.

According to a Gaussian error propagation, 30nA2 = σ2 + (2σ)2 with σ = 13nA

# Endnotes

If unanswered questions occur nevertheless, you can contact me via email (twehrli@student.ethz.ch).

# Text boxes

anti-symmetric BPM deflection [mm]

anti-symmetric BPM deflection [mm]

XBPM reading ["mm"]

XBPM reading ["mm"]

deflection from the reference [mm*]

place in the storage ring [m]

b4

Sum signal [μA]

difference fit-data [„mm“]

c [„mm“]

IB [mm]

IB [mm]

IB [mm]

SLS-TME-TA-2006-0291

April 5, 2006

Correlation between staggered BPM and  RF-BPM readings:

Calibration of their (linear) dependence and reliability of it

Thomas Wehrli

ETH/PSI

PAUL SCHERRER INSTITUT

b3

b2

b1

IB [mm]

IB [mm]

Difference fit-data [„mm“]

linear fit - data

1st BPM, horizonzal deflection [mm]

1st BPM, horizonzal deflection [mm]

Difference Δ1st - Δ2nd [µm]

Sum signal [μA]

blade 2

blade 1

blade current [μA]

time [s]

blade 3

blade 4

BPM bump [mm]

Δ2nd XBPM [µm*]

Δ1st XBPM [µm*]

Sum signal [μA]

2nd BPM, horizonzal deflection [mm]

I B [mm]

XBPM reading ["mm"]

anti-symmetric bumps

symmetric

difference plots ["mm"]

difference plots ["mm"]

XBPM reading [„mm“]

Number of data points

Difference current [nA]

Current blade 1 [μA]

Current blade 2 [μA]

start

end

linear fit - data

linear fit - data

linear fit - data

linear fit - data

cubic fit - data

cubic fit - data

cubic fit - data

cubic fit - data

symmetric bumps

1st XBPM

Screw 1

Screw 2

Screw 3

Sum signal [μA]

linear fit - data

linear fit - data

linear fit - data

cubic fit - data

cubic fit - data

cubic fit - data

cubic fit - data

symmetric bumps

2nd XBPM

anti-symmetric bumps

XBPM reading ["mm"]

I B [mm]